## Slide 1
Manufactured by National Patent Analytical Systems (NPAS), design purchased by Intoximeters, Inc.

## Slide 2: External Features
Heated breath hose
Prevents condensation of water vapor from the breath sample.
Dry gas compartment with internal regulator
Dry gas system delivers a known amount of ethanol to the DataMaster both before and after the subject sample.
Ensures the DataMaster is accurately recognizing and quantitating ethanol concentrations.
Internal regulator controls gas flow into the instrument.
Black coated aluminum cover and RFI antenna
Coating shields against radio frequencies interfering with the internal electronics
Antenna detects radio frequency and stops test sequence

## Slide 3: External Features
USB ports (all the same)
Printer and keyboard connections
Modem port
Downloading and uploading information to and from the instrument
Printer
External laser jet printer

## Slide 4: Internal Features
Infrared Detector and Source
Uses infrared spectroscopy to measure the concentration of ethanol in breath samples for forensic purposes.
Heated Sample Chamber
Prevent condensation of subject sample
Three Filter System
Quantitate ethanol and detect potential interfering substances
Acetone, isopropyl alcohol (sanitizer), methanol (wood alcohol)
Mass airflow sensor
Ensure minimum breath flow rate and volume of breath sample
Internal quartz standard
Known value to check instrument performance

## Slide 5: Software Features
Memory
Holds the most current 2000 tests in memory
Past tests may be reprinted
Secure
Three levels of password protection
Breath Test Operators: Operator
Breath Test Supervisors
Additional menu options
Forensic Scientists
Password for performing certifications, calibrations, etc.

## Slide 6: Home Screen Features

## Slide 7: Home Screen Features
Home Screen Features
Serial Number
Date and Time
Lot Number (External Standard)
Expiration Date (External Standard)
Status Bar (Ready <Push Run>
Barometric Pressure
PSI (External Standard)

## Slide 8: Home Screen Features
DMT Icon
Tank Change
Run
Copy

## Slide 9: Home Screen Features

## Slide 10: Home Screen Features

## Slide 11: Home Screen Features
DMT Icon
  Setup
  Functions
  Reports
  Subject Test – Press to run a subject test
  Diagnostic - Press to run a diagnostic test
  Technician Mode - Shows technician screen with current voltages
  VOC Test - Initiates a Verification of Calibration
  Tank Change – Initiates a tank change to replace the external standard
  Security – Log off or enter your password
  Help – Information about the DMT (current software version)

## Slide 12: Home Screen Features
Setup
  Subject Test Options
    Ask Questions: Allows breath test supervisor to turn off data entry portion at the beginning of the subject test. Default: Yes. Only change for training purposes.
    Number of Tests: Allows breath test supervisor to select number of subject tests in a single breath test sequence. Default: 1. Only change for training purposes.
    Alcohol Display: Turns on and off real-time display of the subject’s ethanol concentration on the bottom of the screen. Default: Yes (On).
    Volume Display: Turns on and off the volume display at the bottom of the screen. Number is the total volume of the subject’s breath sample from the current blow which is compared to minimum volume. Default: No (Off).
    Graph Display: Turns on and off display of the flow and alcohol graph on the screen when the subject has finished providing a breath sample. NOTE: This is not displayed while the subject is blowing, only after their sample is completed. Default: Yes (On).

## Slide 13: Home Screen Features
Setup
  Subject Test Options
    Graph Alcohol: Allows the breath test supervisor to display a real-time graph of subject’s alcohol concentration on the screen while the subject is blowing. Default: No (Off).
    Graph Flow: Allows the breath test supervisor to display a real-time graph of the subject’s breath flow rate on the screen while the subject is blowing. Default: No (Off).
    Query Refusal: Allows the breath test supervisor to turn on a question immediately prior to the subject breath sample to ask if the subject refused. If operator selects Yes, subject result on the test strip will be Refused. If No is selected, subject test will proceed normally. Default: No (Off).
    Wait Between Tests: Option not in use at this time. Default: 0.
    Signature On: Turns on/off the signature option at the end of a subject test sequence. When turned on, the breath test operator will be asked to sign on the touch screen to accept the subject test. Signature will print at the bottom of the subject test strip. Default: Yes (On).
    Copies: Determines how many copies of the test strip are printed after test completion. Default: 1.

## Slide 14: Home Screen Features
Functions
    Set Date/Time
    Remove/Return DMT to service
    Purge Sample Chamber
    Import/Export
      Update from USB
      Output from USB
    Reset Options
    Filter Test – Cycles each filter through the path length 3 times

## Slide 15: Home Screen Features
Setting the Date and Time

## Slide 16: Home Screen Features
Reports
  Last Ticket
  Options – Show current settings
  Breath Records*
  Diagnostic*
  Supervisor Records*
  DMT Update Records*
  VOC Records*
  *Organized by date and time

## Slide 17: Home Screen Features

## Slide 18: Home Screen Features
Technician Mode
  Shows current voltages and temperatures
  Allows some troubleshooting functions such as turning on and off the pump manually

## Slide 19: Subject Tests

## Slide 20: Topics
Pretest Procedures
Alaska Administrative Code
Instrument Preparation
Subject Testing Sequence
Sample Acceptance Parameters

## Slide 21: Subject Test
Breath Test Operators and Breath Test Supervisors should follow proper agency and breath test program procedures when administering a breath test. The subject should be carefully observed during the test.

## Slide 22: Pretest Procedures
The required 15-minute wait must be observed.
Radios should be turned off in the vicinity when possible.
Radio signals may cause a Radio Frequency Interference status message to occur. This will immediately abort your test.
Keep testing area free of contaminants
Solvents, cleaners, mop buckets, exhaust, fumes
Keep subject away from the instrument except when they are providing a sample.
Use a new mouthpiece for each subject test.

## Slide 23: Title 13 Alaska Administrative Code
13 AAC 63.040 Procedure for breath test analysis
(a) The following procedure must be used to obtain and analyze a breath sample on a breath test instrument:
(1) observe the person to be tested for at least 15 minutes immediately before testing, to ensure that the person does not regurgitate or place anything in his or her mouth during that period;
(2) respond to the visual display on the instrument by entering the data requested;

## Slide 24: Title 13 Alaska Administrative Code
(3) when the visual display indicates that the instrument is ready to accept the person's breath sample, instruct the person to blow into the mouthpiece until the visual display indicates that a satisfactory sample has been obtained.
(b) Only a person certified as an operator under 13 AAC 63.050 may operate a breath test instrument, except that a supervisor may supervise the use of an instrument by a non-certified person for training or demonstration purposes.

## Slide 25: Instrument Preparation
The instrument status box on the bottom left hand of the display screen should read “Ready <Push Run>”.
Verify the time and date are accurate in the top right corner of the screen.
Be prepared to supply a new mouthpiece for each testing sequence.

## Slide 26
Date and Time
103 PSI

## Slide 27: Setting Date and Time
Verify the date and time are correct before running a subject test.
If they are not:
  Press the DMT Icon in the top left corner
  Press the Functions menu option
  Press Set Date/Time
  Enter the correct date and time
  Press OK to accept

## Slide 28
103 PSI

## Slide 29
103 PSI

## Slide 30: Subject Testing Sequence Data Entry
Press the RUN button to initiate the test.
Enter the operator password.
Choose No for scan driver's license and type license information.
Select the test type and enter the required information for subject, operator, and arrest information.
  If the test type is “Test” the subject name will automatically populate with ‘Test, Test’. This is typically used for training or demonstration purposes and is not used for evidential breath tests.

## Slide 31
Push Run
103 PSI
Approved and Issued: 4/7/2017 CRF

## Slide 32
OPERATOR
103 PSI
Approved and Issued: 4/7/2017 CRF

## Slide 33: Subject Testing Sequence Data Entry
Test types
  DUI
  Minor Consuming
  Other
  Test
Subject Name – enter the full name of subject being tested

## Slide 35: Subject Testing Sequence Data Entry
OLN - Enter the driver’s license number of subject being tested
Operator Name
Operator Number
Dept/Agency
Case/Report

## Slide 37: Subject Testing Sequence Breath Analysis Portion
The breath analysis portion of the test is automatic. Once in progress, the operator cannot alter the outcome of the test.
The duty of the breath test operator during the breath analysis portion of the sequence is to instruct the subject how and when to provide a sample.

## Slide 38: Subject Testing Sequence Breath Analysis Portion
As the subject is providing a sample, the instrument will give a sustained tone if the sample is above the minimum flow rate.
If the sample is below the minimum flow rate, the instrument will give an intermittent tone (beeping).

## Slide 39: Subject Testing Sequence
Diagnostic Check
Purging and Ambient Air Testing
Blank Test
Internal Standard Check
External Standard #1
Purge, Blank Test
Question: Did the subject refuse? (Only if Refusal Query is On)
PLEASE BLOW
Purge, Blank Test
External Standard #2
Purge, Blank Test
Data saved
Test strip prints automatically

## Slide 40: Subject Testing Sequence Breath Analysis Portion
Signature Box
  If your agency has the signature option turned On, a signature box will appear at the end of the subject test sequence. You can sign using the touch screen and the signature will appear at the bottom of the test strip.

## Slide 41: Example of a Subject Test Printout
Breath flow rate (liters/minute)
Time in seconds
Minimum flow rate
Subject breath profile (flow rate)
Breath sample volume (liters)

## Slide 42: Graph
There is no scale for the alcohol level. This is purposeful.
A breath sample must meet ALL the sample acceptance requirements to produce an alcohol level result.
The dotted line is the subject’s flow rate in liters/minute (L/min).
The solid line is the subject’s alcohol level in grams ethanol/210 liters of breath (no alcohol measured on previous slide).
The solid line at 3 liters/minute is the minimum flow rate required.

## Slide 43: Sample Acceptance Parameters
Minimum flow rate of 3.0 L/min
Minimum breath volume of approx. 1.5 L
The breath alcohol curve must have reached a plateau (defined by the manufacturer).
The breath alcohol curve must not trigger an invalid status message.
10 or less attempts within the 2-minute window.

## Slide 44: Completed Test
Once the subject has provided a breath sample that meets the acceptance parameters, the instrument will continue with the second external standard and final blank test.
Signature Box (optional)
Test strip is printed automatically.

## Slide 45: Completed Test
IMPORTANT: It is the responsibility of the breath test operator to ensure that the breath test printout contains a completed breath test sequence.

## Slide 46: Non-Completed Test
If a status messages occurs after the subject sample, the test will abort at that point and the breath result will be reported as “XXX”.
The DataMaster will list the status message and prompts you to restart the test.
Further information is provided in the Status Messages presentation of this course.

## Slide 47: Non-Completed Test

## Slide 48: Subject Testing Sequence Breath Analysis Portion
Hands on Exercise
  Perform a non-drinking subject test
  Perform a non-completed subject test (trigger any status message)
    Incomplete
    Maximum Attempts Exceeded

## Slide 49: Tank Changes

## Slide 50
Dry Gas External Standards
DataMaster DMT uses commercially-prepared dry gas cylinders.
  Mixture of Nitrogen gas / Ethanol
  Provided by the Crime Lab

## Slide 51: Dry Gas External Standards
The instrument tests the dry gas external standard twice in a subject testing sequence.
  Once before the subject sample and once after
Demonstrates the instrument is accurately measuring ethanol at the 0.080 level.

## Slide 52: Dry Gas External Standards
Internal regulator (no need to turn On or Off)
Every tank will have:
  Target value of 0.080
  Lot number and Expiration date
Expiration Date
Lot Number

## Slide 53: Dry Gas External Standards
Lot number, expiration date, and target value print on the subject test strip. This information comes from the tank change.
The DataMaster DMT informs users when the dry gas standard is less than 150 psi (warning) or reaching the expiration date of the tank.
The DataMaster will not run a test if the PSI in the tank is less than 100 or the tank is expired.

## Slide 54: Dry Gas External Standards
Tank PSI
Prepopulated
103 PSI

## Slide 56
103 PSI
Turns red if less than 150 PSI (warning level)
Approved and Issued: 4/7/2017 CRF

## Slide 57
103 PSI
Press to start Tank Change
Approved and Issued: 4/7/2017 CRF

## Slide 58
Remove old tank
(turn counter-clockwise)

## Slide 59: Changing the External Standard
Enter your operator name and number.
Enter the lot number and expiration date from the tank label.
  NOTE: Ensure information is read directly from the tank and entered correctly.
This information will be requested a second time and if it does not match the initial entry, you will get a message that the information doesn’t match.
The DMT will prompt you to ensure the tank is in place in the compartment and it will then perform a VOC.

## Slide 60
AG608504
03/25/2018
From tank label

## Slide 61
Tank information is requested a second time
ENTRIES MUST MATCH

## Slide 63
Insert new tank
(turn clockwise)

## Slide 65: DataMaster DMT will automatically initiate a VOC. This will take a few minutes to complete.

## Slide 66: Changing the External Standard
A subject test cannot be performed until a VOC has been successfully completed.
Breath Test Supervisor will demonstrate a tank change for the class.

## Slide 67: Changing the External Standard
Breath Test Operators and Supervisors can both change the tank.
The agency breath test supervisor requests tanks from the Crime Lab when running low or approaching expiration.

## Slide 68: Review
When are VOCs initiated?
1. Required after a tank change
2. Automatically initiated by the instrument on VOC days
3. When the VOC Test option is selected

## Slide 69: Review
How many tests of the external standard are done during a subject test?
2

## Slide 70: Review
What are the minimum volume AND flow rate requirements for a subject sample?
~1.5 L of air
3.0 L/Min

## Slide 71: Review
What two things are graphed on the subject test printout?
The flow rate (dashed line)
Alcohol level (solid line)