## Slide 1: IR Theory (How the DataMaster Works)
Measurement Definitions
Absorption of Infrared (IR) Light
Analytical Components of DataMaster

## Slide 2: Terminology of Instrumental Analysis
An instrument measures some property of a sample.
Examples:
A thermometer measures the temperature of the air.
A speedometer measures the speed of your car.
A scale measures the weight of the apples you purchased.
A DataMaster measures the ethanol concentration in a breath sample.

## Slide 3: Terminology of Instrumental Analysis
Adjustment: The process of creating a relationship between signal and a known amount (ex: setting the dial on your scale)
Calibrator: The known amount used to create a relationship with instrument signal.
Calibration: Demonstrating the relationship of signal and concentration using other known amounts.

## Slide 4: Terminology of Instrumental Analysis
Examples
Unknown Sample: Weighing yourself on the scale
Control: Weighing a 50 lb bucket of sand to see what the scale reads
Unknown Sample: A subject’s exhaled breath
Control: A sample of known concentration used to check the calibration of the instrument

## Slide 5: Terminology of Instrumental Analysis
Alaska DataMaster ADJUSTMENT may occur at
a Manufacturer facility OR
the Alaska State Crime Lab
ADJUSTMENT is NOT done in the field.
DataMasters are checked in the field (more on this shortly).

## Slide 6: Terms Used in this Class
Instrument synonyms
DataMaster DMT, DataMaster, Breath Instrument, DMT, Evidentiary Breath Test
Samples
Breath Sample, Subject Sample
Controls
Internal Standard, Quartz Plate
External Standard, Dry Gas Standard
Blank Test, Negative Control
Wet Bath Simulators

## Slide 7: Terms Used in this Class
Quality Assurance (QA)/Quality Control (QC)
  VOC = Verification of Calibration
  Instrument Calibration process
    Performed by the Crime Lab

## Slide 8: Analytical Principle of DataMaster
What signal does the DataMaster use to determine breath alcohol concentration?
IR Energy
IR Energy
Vehicle Window Tint
Fog decreasing visibility of mountains
Absorption of infrared light
More light absorbed = higher concentration of ethanol

## Slide 9: Infrared Light

## Slide 10: Infrared Light
Infrared is emitted from heated objects
The human eye cannot see infrared light

## Slide 11: Concept Drawing of DataMaster
Parts described:
  Breath Hose, Exhaust, Detector, Sample Chamber, Source

## Slide 12: Infrared Spectroscopy
Chemicals (like ethanol) absorb infrared (IR) light based on their structure
How do chemicals differ from one another?
Could be made up of different elements
Could be arranged in different ways
This leads to changes in chemical properties
Infrared light causes chemical molecules to vibrate in a way that unique to each different chemical.

## Slide 13: Infrared Spectroscopy

## Slide 14: Structure of Ethyl Alcohol (Ethanol)
Carbon – Carbon bond (1)
Carbon – Oxygen bond (1)
Carbon – Hydrogen bonds (5)
Oxygen – Hydrogen bond (1)
C
C
H
H
H
H
H
H
O

## Slide 15: Infrared Spectroscopy
Only certain wavelengths of IR are absorbed by a molecule. These wavelengths are then missing after the IR reaches the detector.
IR Transmittance through ethanol is reduced to less than 50% at wavelengths of interest
3.37
3.44
3.50
An infrared spectrum can be thought of as a chemical’s unique fingerprint.

## Slide 16: The Colors We See are NOT Absorbed

## Slide 17: Infrared Spectroscopy
How the DataMaster ensures the reported concentration is from ethanol only:
  Breath sample’s absorption pattern at 3 wavelengths is used to ensure that the concentration reported is from ethanol and not something else.
  Wheel with three filters (3.37, 3.44, 3.50 mm) located between sample chamber and detector
  If the pattern does not match what is expected, INTERFERENCE DETECTED status message occurs.

## Slide 18: The Optical System
The Optical System consists of the following components:
IR source
Sample chamber
Mirrors
Windows
Focusing lens
Quartz internal standard
Chopper wheel
Three IR filters
IR detector

## Slide 19: IR Source – Where the light comes from

## Slide 20: Sample Chamber
Sample chamber (holds the breath sample being analyzed) is made of highly polished aluminum.
  Breath Hose and Sample Chamber are heated to 50oC to prevent breath condensation
  Interior is folded back on itself to increase sensitivity.
  Sample Chamber volume is very small when compared to minimum breath volume requirement.

## Slide 21: Mirrors
Mirrors are located at the “folds” of the sample chamber and designed to reflect IR energy completely.

## Slide 22: Quartz Windows

## Slide 23: Focusing Lens
The focusing lens is located where the IR light leaves the sample chamber.
The lens gathers the IR energy into one focused beam directed onto the detector.

## Slide 24: Three Filter System
The filters block IR light from entering the detector EXCEPT the critical wavelengths. The purpose is to ensure results are specific to ethanol.
Absorbance at 3 wavelengths used:
3.37 m
3.44 m
3.50 m

## Slide 25: Internal Standard
Quartz Internal Standard/Quartz Plate
Checks that the programmed calibration factors are accurate.
Filters & quartz plate are embedded in two geared wheels located between the sample chamber and detector.

## Slide 26: Filter and Internal Standard Wheels
Quartz Plate
(Other 3 hole empty)
Filter 1
Filter 2
Filter 3
Rotation allows different combinations of filters

## Slide 27: Detector
Absorbs remaining IR light from 3.37 m, 3.44 m, 3.50 m and converts the light into an electrical signal.

## Slide 28: Summary
Why Infrared Light?
  Alcohol molecules absorb infrared light at specific wavelengths that other common breath gases do not. This makes infrared testing very good at detecting alcohol without getting confused by other substances.