## Slide 1: Status Messages and Troubleshooting

## Slide 2: What Is a Status Message?
Usually does not indicate the instrument is broken.
Often indicates an issue with the sample being provided or the environmental conditions of the room.

## Slide 3: Common Status Messages
Incomplete
Invalid
Interference Detected
Blank Error
Ambient Fail
Radio Frequency Detected (RFI)
Standard Out of Range
Maximum Attempts Exceeded
Pump Error

## Slide 5: Incomplete
The subject sample did not meet the minimum requirements within the two-minute window:
  Minimum Flow: 3 L/min
  Minimum Volume: ~1.5 L
  Breath Alcohol Plateau
    (Memorize all 3!)
The subject flow must fall below minimum flow for the sample to be accepted. The instrument will NOT cut the subject off when minimum requirements are met (allows the person to keep blowing).
Requirements start over each time the person stops blowing – aka one “attempt”. They are not cumulative.

## Slide 6: Incomplete

## Slide 7: Incomplete
Why minimum requirements?
To ensure that we are sampling deep lung air - this is most closely related to what is in the blood.

## Slide 8: Invalid
Indicates the subject’s breath profile contained a negative slope.
Means the alcohol concentration went up and then down – not typical
Datamaster software determines this - Cannot be reliably determined by looking at the graph
Possible Causes
Mouth alcohol
Improper blowing

## Slide 9: “Ideal” Breath Profile

## Slide 10: Should I Restart the 15 Minute Wait?
If the operator believes that the subject
1) Placed anything in their mouth OR
2) Brought stomach contents into their mouth (regurgitated)…
They should restart the 15-minute wait.
If no reason to believe stomach contents came up, it is up to the individual agency policy.
Why??? – If the stomach contained alcohol, that alcohol may now be in the mouth. Alcohol in the mouth could trigger….What?
Invalid status message

## Slide 11: Interference Detected
The instrument is indicating the presence of a substance other than ethanol in the subject sample.
DataMaster uses three IR filters to look for interfering substances.
If the pattern of the three filters looks different than the saved pattern, Interference Detected is triggered.

## Slide 12: Interference Detected
What cause interference?
Acetone (diabetics, possibly keto diet)
Huffing (inhaling gases – e.g., Dust Off)
What should you do?
Run the breath test again
If interference occurs a second time, get a search warrant for blood.
Inhalants leave the body quickly!

## Slide 13: Interference Detected
If interference detected, test the person two times at most (more tests likely to produce the same result).
Do NOT move to another DataMaster.
If this happens, obtain a legal blood sample.

## Slide 14: What does Interference Detected indicate?
Something other than ethanol in the subject’s breath
NOTE: Could be a combination of ethanol with another substance

## Slide 16: Blank Error
Indicates the instrument is unable to get a zero baseline.
Usually caused by environmental factors:
Ethanol or fumes in the room air such as:
  Cleaning products
  Vehicle exhaust
  Heavy dust

## Slide 17: Blank Error: What do I do?
If the subject smells of alcohol or chemicals, move them away from the instrument except for while providing their sample.
If the room smells of cleaning products, fumes, or is very hot, ventilate the room and try again.
Ensure the back of the instrument is free from obstructions (stacks of papers, etc.). This is where the instrument ventilates.

## Slide 19: Ambient Fail
Like Blank Error but occurs at the beginning of the test sequence.
Follow the same troubleshooting as Blank Error.

## Slide 20: Radio Frequency Interference (RFI) Detected
Indicates the instrument has detected radio frequency in the vicinity.
Black instrument coating shields the instrument from RFI.
RFI antenna is inside the breath hose.
Antenna detects the radio frequency and triggers the instrument to abort the test.

## Slide 21: RFI Detected
Does not mean the instrument was affected by RFI, just that its was detected. The instrument stops the test as a precaution.
If you receive this status message, ensure all radios in the area are turned Off. Then restart the test.
If you cannot turn Off your radio, move the radio as far away as possible or turn of transmitting or you will trigger another RFI status message.

## Slide 23: Standard Out of Range
Refers to the external standard (dry gas tank) value
Measurement was outside of the allowable range
Allowed range is +/- 0.005 of the target value
Target value is adjusted for barometric pressure

## Slide 24: Standard Out of Range
The dry gas cylinder used by the instruments for Quality Assurance have a known value (the target).
With each subject test, the cylinder is analyzed twice (before and after the subject blows).
Both readings must fall within the +/- 0.005 of the target value adjusted for barometric pressure.
If not, the test is aborted.

## Slide 25: Standard Out of Range: What do I do?
Restart the breath test.
If the status message occurs again, move to another instrument (if possible) OR get a legal blood sample.
Notify your Breath Test Supervisor.
If your agency does not have a Breath Test Supervisor (uncommon), contact the Crime Lab at dps.scdl.toxicology@Alaska.gov.

## Slide 26: Maximum Attempts Exceeded
The subject provided 10 or more “attempts” during the two-minute window.
An “attempt” is when the subject’s breath flow first exceeds and then drops below the minimum flow rate (started blowing and then stopped).
On the 10th “attempt”, if the subject stops blowing without meeting the sample acceptance criteria, this status message will occur.

## Slide 27: Exceed attempts

## Slide 28: Other Status Messages
The status messages discussed previously are the most encountered.
For most status messages, running the test a second time is good practice.
A Troubleshooting Guide is located at the end of the Alaska Breath Test Operator and Supervisor Manual to help you troubleshoot other problems.
Datamaster - DataMaster - CrimeLab - Alaska Department of Public Safety
If you experience a problem, always notify the Breath Test Supervisor, even if you fix it.
Save all documentation.

## Slide 29: Status Message Dialog Box

## Slide 30: Completed Tests
Only completed tests have a breath # displayed on the test strip.
If the # is present, no status messages occurred during the breath test sequence.

## Slide 31: Completed Test Strip

## Slide 32: Non-Completed Tests
All tests stop when a status message occurs.
If the status message occurs after the subject sample (blow), the breath value will be listed as “XXX” on the test strip (i.e., breath test was not completed).
The Crime Lab will not testify to non-completed tests.

## Slide 33: Non-Completed Test Strip

## Slide 34: Other Problems
License/Card reader not working
Printing not working
Instrument display frozen

## Slide 35: License/Card Reader Not Working
The Crime Lab does not maintain the card readers on the DataMaster DMT.
If the card reader does not work, manually input the identifying information.

## Slide 36: Printing Problems
If your test strip does not print…
  Make sure the printer is On, plugged in, has paper, and does not having any warning lights.
  Press the COPY button on the home screen to print the last test in memory.
  If no test is printed, unplug the printer USB cable from the DataMaster. Wait 5 seconds, plug back in, then wait for the driver error popup. Click the X in the top right-hand corner of the popup window (Do not hit OK).
  Hit Copy again. If the test strip does not print, press the DMT icon in the upper left-hand corner of the touch screen, then select Reports, then Breath Records on the right-hand screen. Verify that your test is listed.
    If the test does not show up in Breath Records, the test did not save.
    If the report does show up in Breath Records but still will not print, contact your Breath Test Supervisor.

## Slide 37: Printing Problems (cont.)
If your test is shown in the Reports screen, it can be printed when the printer is working again.
If your test does not show in reports, run the test again.
Contact your Breath Test Supervisor for more assistance.

## Slide 38: Instrument Touch Screen Frozen
Reboot the DataMaster using the black toggle switch on the back.
If this does not solve the problem, contact your Breath Test Supervisor.

## Slide 39: Reminders
Keep your Breath Test Supervisor informed of problems with the instrument. Let them know as soon as possible, even if you fixed the problem.
Save any printed documentation or status messages that print.
Check your test strips after the test completes to ensure they are complete and contain no status messages.

## Slide 40: Review
What does an INCOMPLETE status message indicate on a subject test strip?
The subject did not provide the minimum sample required within the two minutes allowed.

## Slide 41: Review
What does an INVALID status message indicate?
The instrument detected a negative slope in the breath profile. This could be caused by mouth alcohol or improper blowing.

## Slide 42: Review
An INTERFERENCE DETECTEDstatus message can be caused by a radio being keyed in the room. (True/False)
False: Interference Detected status messages indicate that something other than just ethanol was detected.

## Slide 43: Review
Cell phones should always be turned off around the DataMaster to prevent RFI DETECTED status messages.
(True/False)
False: Cell phones generally do not trigger RFI status messages. Operators should ensure radios are turned Off in the vicinity of the instrument when possible to prevent RFI status messages.

## Slide 44: Review
What status message could occur if a subject who smelled strongly of alcohol was too close to the instrument during the test sequence?
BLANK ERROR or AMBIENT FAIL

## Slide 45: Review
What does a STANDARD OUT OF RANGE status message indicate?
The external standard test value was outside of the allowable range.
(+/- 0.005 of the target value adjusted for barometric pressure)

## Slide 46: Any Questions? (Exam is the next thing)

## Slide 47: Reminder: Exam format
Required: 75% or higher to pass
Closed book exam completed independently without the use of course material.
Electronic Examination via ACADIS Portal
Special needs or accommodations - Speak with an instructor now